## Circuit
[6 empty] | General Circuit |  | *O2 too far from patient"
[8 empty] | Capped Bleed Port |  |  |  |  | O2 Tubing | Viral Filter
[6 empty] | ST Machine | 22mm ID Adaptor | Standalone Exhalation Port | Viral Filter | CPAP Tubing |  |  | O2 Bleed in Adaptor | Exhalation Port | Spirometry / Gas Sampling (Optional) | Patient
[8 empty] | *Must have viral filter attached to CPAP tubing* Provides filter for all patient exhalations | [5 empty] | Capped Bleed Port
[13 empty] | *Cannot use adaptor on exhalation port - O2 exits port Use 90 degree O2 bleed-in adaptor, 1 position removed from gas sampler* |  |  |  | Part Counts
[17 empty] | VPAP ST | 22x22 Adaptor | F&P RT017 / Phillips 1065775 | CPAP Tubing | F&P RT 219 | Viral Filter
[6 empty] | Equivalent Circuit 1 | [6 empty] | O2 Tubing | Intersurgical 441-1844030 | Spirometry / Gas Sampling (Optional) | Patient | 1 | 1 | 1 | 0 | 1 | 2
Capped Bleed Port
[6 empty] | ST Machine | 22mm ID Adaptor | Phillips 1065775 Exhalation Port | Intersurgical 441-1844030 | Fisher & Paykel RT219 OR Standard Tubing (Tubing) |  |  | Teleflex 1642 | F&P RT219 (Exhalation Port)
[6 empty] | Equivalent Circuit 2 | [6 empty] | O2 Tubing | Intersurgical 441-1844030 | Spirometry / Gas Sampling (Optional) | Patient
[8 empty] | Capped Bleed Port | [8 empty] | 1 | 1 | 2 | 1 | 0 | 2
[6 empty] | ST Machine | 22mm ID Adaptor | Phillips 1065775 Exhalation Port | Intersurgical 441-1844030 | Standard CPAP Tubing |  |  | Teleflex 1642 | Philips 1065775
| Product (Links) | Description | Optional? | Critcial Aspects | Function | Comparable Products (Links) |  | ST Machines | Sizing Adaptors | Exhalation Port Adaptors | Filter | Breathing Circuit Tubing | CPAP Standard Tubing | O2 Adaptors | O2 Tubing | Gas Sampler | HME
1 | Philips DreamStation | Bilevel Machine |  | Bilevel capability Min pressure of 25 cm H2O | Generates pressure | ST Machines |  | Resmed S9 VPAP ST | Teleflex 1421 | Phillips 1065775 | Intersurgical 441-1844030 | Fisher & Paykel RT219 | Phillips 22mm Tubing | Vyaire 004081 | Medline HCS4524G | GE 889560 | Medline HCSMC750
2 | Teleflex 1421 | 22mm X 22mm Adaptor |  | 22mm X 22mm | Accepts RT017 Adaptor | Sizing Adaptor |  | Phillips DreamStation ST30 |  | Fisher & Paykel RT017 | Intersurgical DYNJAABV1 |  | ResMed 22mm Tubing | Vyaire 004204 | Medline HCSU4524 | GE 8004382 | Medline DYNJAAHME9
3 = 8 | Philips 1065775 | Exhalation Filter |  | Exhalation | Provides pressure exhalation at critical location | Exhalation Port Adaptors |  | Phillips System One |  | *Only available as kit RT219, standard tubing + RT017 adapter* | Vyaire 001851 |  |  | Teleflex 1642 | Teleflex 1925 |  | Vyaire M1038637
4 = 9 | Intersurgical 441-1844030 | Viral Filter |  | 22 Outer Diameter 15 Inner Diameter | Filters blower of Bilevel machine AND exhalation port to prevent release of particles. | Viral Filters |  |  |  |  | Hudson RCI 1605 |  |  | Resmed 26960 | Vyaire 1350 |  | Vyaire 557071500
5 | Fisher Paykel RT219 | Breathing Circuit Tubing + Exhalation Port | Can use standard tubing with exhalation port / CAPPED O2 adaptor | Standard End + Exhalation Port "Leak Proof Tubing" | Carries air to pt, exhalation port | Breathing Circuits | [10 empty] | Teleflex 19912
5a | Standard CPAP Tubing | CPAP Tubing |  | "Leak Proof Tubing" | Carries air to pt, exhalation port | CPAP Tubing | [10 empty] | Fisher Paykel MR-850-JITA
6 | Teleflex 1642 | O2 Bleed-in Adaptor |  | O2 bleed valve not directly across from exhalation port | O2 Input | O2 Adaptors
7 | Medline HCS4524G | O2 Tubing |  | Universal Connector O2 tree and O2 Connector Set Length, 7' - 14' | O2 Input | O2 Tubing
8 = 3 | Philips 1065775 | Exhalation Filter | Can use tubing with attached port | Filtered Exhalation | Provides pressure exhalation at critical location | Exhalation Port Adaptors
9 = 4 | Intersurgical 441-1844030 | Viral Filter |  | 22 Outer Diameter 15 Inner Diameter | Filters blower of Bilevel machine AND exhalation port to prevent release of particles. | Viral Filters
10 | GE 8004382 | Gas Sampler | Yes |  | Allows sampling of exhalations | Gas Sampler
11 | Medline HCSMC750 | HME | Can replace with in-line humidifier | Dead space under 60ml Moisture Output 30mg | Humidifies circuit for pt comfort | HME
12 | GE | Control Module | Yes, monitoring | CO2 Monitoring FiO2 Tidal Volume Monitoring | Monitors

## Components
| Product (Links) | Description | Optional? | Critcial Aspects | Function | Comparable Products (Links) | ST Machines | Sizing Adaptors | Exhalation Port Adaptors | Filter | Breathing Circuit Tubing | CPAP Standard Tubing | O2 Adaptors | O2 Tubing | Gas Sampler | HME
[7 empty] | Resmed S9 VPAP ST | Teleflex 1421 | Phillips 1065775 | Intersurgical 441-1844030 | Fisher & Paykel RT219 | Phillips 22mm Tubing | Vyaire 004081 | Medline HCS4524G | GE 889560 | Medline HCSMC750
1 | Philips DreamStation | Bilevel Machine |  | Bilevel capability Min pressure of 25 cm H2O | Generates pressure | ST Machines | Phillips DreamStation ST30 |  | Fisher & Paykel RT017 | Intersurgical DYNJAABV1 |  | ResMed 22mm Tubing | Vyaire 004204 | Medline HCSU4524 | GE 8004382 | Medline DYNJAAHME9
2 | Teleflex 1421 | 22mm X 22mm Adaptor |  | 22mm X 22mm | Accepts RT017 Adaptor | Sizing Adaptor | Phillips System One |  | *Only available as kit RT219, standard tubing + RT017 adapter* | Vyaire 001851 |  |  | Teleflex 1642 | Teleflex 1925 |  | Vyaire M1038637
3 = 8 | Philips 1065775 | Exhalation Filter |  | Exhalation | Provides pressure exhalation at critical location | Exhalation Port Adaptors |  |  |  | Hudson RCI 1605 |  |  | Resmed 26960 | Vyaire 1350 |  | Vyaire 557071500
4 = 9 | Intersurgical 441-1844030 | Viral Filter |  | 22 Outer Diameter 15 Inner Diameter | Filters blower of Bilevel machine AND exhalation port to prevent release of particles. | Viral Filters | [9 empty] | Teleflex 19912
5 | Fisher Paykel RT219 | Breathing Circuit Tubing + Exhalation Port | Can use standard tubing with exhalation port / CAPPED O2 adaptor | Standard End + Exhalation Port "Leak Proof Tubing" | Carries air to pt, exhalation port | Breathing Circuits | [9 empty] | Fisher Paykel MR-850-JITA
5a | Standard CPAP Tubing | CPAP Tubing |  | "Leak Proof Tubing" | Carries air to pt, exhalation port | CPAP Tubing
6 | Teleflex 1642 | O2 Bleed-in Adaptor |  | O2 bleed valve not directly across from exhalation port | O2 Input | O2 Adaptors |  |  |  | Equivalents
7 | Medline HCS4524G | O2 Tubing |  | Universal Connector O2 tree and O2 Connector Set Length, 7' - 14' | O2 Input | O2 Tubing
8 = 3 | Philips 1065775 | Exhalation Filter | Can use tubing with attached port | Filtered Exhalation | Provides pressure exhalation at critical location | Exhalation Port Adaptors
9 = 4 | Intersurgical 441-1844030 | Viral Filter |  | 22 Outer Diameter 15 Inner Diameter | Filters blower of Bilevel machine AND exhalation port to prevent release of particles. | Viral Filters
10 | GE 8004382 | Gas Sampler |  |  | Allows sampling of exhalations | Gas Sampler
11 | Medline HCSMC750 | HME | Can replace with in-line humidifier | Dead space under 60ml Moisture Output 30mg | Humidifies circuit for pt comfort | HME
12 | GE | Control Module | Yes, monitoring | CO2 Monitoring FiO2 Tidal Volume Monitoring | Monitors

## Circuit Builds
General Circuit |  | *O2 too far from patient* for effective use
|  | Capped Bleed Port |  |  | O2 Tubing | Viral Filter
ST Machine | 22mm ID Adaptor | Standalone Exhalation Port | Viral Filter | CPAP Tubing | O2 Bleed in Adaptor | Exhalation Port | Spirometry / Gas Sampling (Optional) | Patient
|  | *Must have viral filter attached to CPAP tubing* Provides filter for all patient exhalations |  |  |  | Capped Bleed Port
*Cannot use adaptor on exhalation port - O2 exits port instead of going to patient Use 90 degree O2 bleed in adaptor, 1 position removed from gas sampler*
Equivalent Circuit 1
|  | Capped Bleed Port |  |  | O2 Tubing | Viral Filter | Spirometry / Gas Sampling (Optional) | Patient
ST Machine | 22mm ID Adaptor | Phillips 1065775 Exhalation Port | Intersurgical 441-1844030 | Fisher & Paykel RT219 OR Standard Tubing (Tubing) | Teleflex 1642 | F&P RT219 Exhalation Port
Equivalent Circuit 2
|  | Capped Bleed Port |  |  | O2 Tubing | Viral Filter | Spirometry / Gas Sampling (Optional) | Patient
ST Machine | 22mm ID Adaptor | F&P RT219 Exhalation Port | Intersurgical 441-1844030 | Standard CPAP Tubing | Teleflex 1642 | Philips 1065775 = F&P RT017